Campus:_______________________			Date:_________

Building:______________________	Room:__________	Hood#:_______

Date of last evaluation:_________	Monitoring Equipment:_________________

CURRENT HOOD USE

List any dedicated uses of this hood:____________________________________

Yes	No

Is the hood routinely used for storage of equipment or chemicals?

Are the slots blocked by objects?

Are processes or equipment located within 6 inches of the sash or rear slots?

Is the bench area inside the hood clean and free of debris and residues?

Is large equipment in use in the hood that could block or reduce air flow?

Are open or unlabeled chemical containers (not in use) present?

Do any large heated processes or equipment affect the air floor in the hood?

Does the sash adjust easily?

Is the sash glass in good condition?

Does the light work?

Is the light cover in place and intact?

Are hood fixtures in working order?

Are hood fixtures corroded or dirty?

Are interior surfaces and baffles in good condition?

Does the hood have a visual indicator of operating status (on/off)?

Are flammable liquids stored below the hood?

Flow Measurement Procedure:

- Indicate in the appropriate space on the next sheet the position of any fixed equipment in the hood.

- Visually divide the hood face (full open sash position) into a grid containing 9-12 equal sections.

- Measure the air velocity at the center of each section and record the velocity in the appropriate location on the diagram on the next page. Determine the average of the recorded data and record the result in the space indicated.

- Use stannic oxychloride “smoke tubes” to observe air flow patterns at the hood face and within the hood. Note any eddies or unusual flow patterns in the grid diagram.

- If the hood does not supply 100 lfpm +/- 10% at full open sash position, determine operating height by lowering sash height until average airflows are 100 lfpm +/- 10%. Record operating height in the space indicated, and mark the operating height on the hood. Sash stops must be installed as soon as possible. Hood maintenance should be scheduled to bring operating velocity to 100 lfpm +/- at full open sash position.

- Attach any recommendations to this form.

Fixed Equipment Diagram:

Full Open Sash Position Velocity Measurement:

Dimensions of opening:___’ H x ___’W=___ sqft   X=___ lfpm Average Flow: ____cfm

Operating Height Velocity Measurement (if Full Open Average Flow < 100 lfpm +/- 10%:

Sash operating height:_____’

Dimensions of Opening: ___’H x ___’W= ___sqft   X=___ lfpm  Average Flow: ____cfm